## X RAY ATTENUATION
Determination of a technique for thin film measurements
ranging from 50nm to 1um(Copper To Kapton)
with a resolution of +-5nm using X-Ray attenuation
AIM-
To determine a technique for thin film measurements
ranging from 50nm to 1um(Copper To Kapton)
with a resolution of +-5nm using X-Ray attenuation
DEVICES REQUIRED-
Si-PIN X-Ray Detector | https://www.amptek.com/products/x-ray-detectors/sipin-x-ray-detectors/sipin-x-ray-detectors
Mini-X2 X-Ray Tube | https://www.amptek.com/products/mini-x2-x-ray-tube
PX2T-CdTe Power Supply & Shaping Amplifier | https://www.amptek.com/internal-products/obsolete-products/px2t-cdte-power-supply-shaping-amplifier-for-the-xr100t-cdte
Ratemeter | https://122.physics.ucdavis.edu/course/cosmology/sites/default/files/files/Gamma%20Spec/PulseHeightAnalysis.pdf
BNC CABLES
RS232
USB PORT
Settings of PX2T-CdTe Power Supply & Shaping Amplifier
RTD | ON
EXPERIMENTAL SETUP
The setup above should be place in a lead box to prevent interference with cosmic radiations
The setup below has the detector located 2.5cm away feom the aperature and the
X ray tube located 5.5 cm away from the aperture
The X ray detector has power pin(Hirose MQ172-3PA(55),
Mating plug: MQ172-3SA-CV),(RS232),and an output pin(BNC CABLE)
Ground
POWER CABLE
MICRO USB PORT TO PC
|  | POWER CABLE | OUTPUT
The power cable of the detector goes to the amplifier PX2T/CR which provides the power supply
for the detector
The output signal of the detector goes to the Single Channel Analyzer
The ouput from SCA goes to the ratemeter which measures the count rate
PROCEDURE
1)
To generate X rays, a Mini-X2 X-Ray Tube is used which includes the X-ray tube,
the power supply, the control electronics, and the USB communications to the computer.
The Mini-X2 consists of two components: the Mini-X2 X-ray tube module and the Mini-X2 Controller.
The X-ray tube module includes the tube and high voltage power supply.
The X ray tube is operated by a software called Mini X Controller
2) | To detect X rays, Si-PIN X-Ray Detector is used .The Si-PIN is a high performance thermoelectrically
cooled X-ray detector and can be combined with several preamplifier and digital pulse processor configurations.
THe signal from the detector goes to the SCA which is connected to a ratemeter
3)
a) | Ratemeter measures the counting rate of randomly arriving pulses with 18 distinct scales that span this range of counting rates.
b) | The scales are set up in a 25, 50, 100 order, ranging from 25 counts per second to 10^7 counts per second full scale.
c) | To read counts per second we have to look at the Front-panel meter.
d) | Actual value for the full-scale reading is determined by the product of the RANGE and and MULTIPLIER values.
e) | The MULTIPLIER switch selects a multiplying factor of 0.5,1.0, or 2.0. and there are three range options, 0-2.5,0-5 and 0-10
f) | So we connect BNC from the output from the detector/amplifier to the input of the SCA(Single Channel Analyzer).
g) | When the peak signal amplitude of a single channel analyzer (SCA) falls within the pulse-height window
defined by two predetermined threshold values, an output logic pulse is generated.
h) | This signal goes to the ratemeter which neasures the count rate of number of peaks of the signal
4) | Now count rates can be observed with change in parameters like tthickness voltage current and gain. We will analyze the count
trate for various thickness with change in parameters
5.5cm
2.5cm

## OBSERVATIONS AND RESULTS
OBSERVATIONS
1) | Table corresponding to counts for values of current and voltage
Counts for Voltage
| Current | 10KV | 20KV | 30KV
| 10 | 400 | 1500 | 3900
| 20 | 1100 | 2700 | 7800
| 30 | 1600 | 4300 | 9600
| 40 | 2200 | 6400 | 10000
| 50 | 2700 | 7400 | 10000
| 60 | 3200 | 8100 | 10000
| 70 |  | 9000
| 80 |  | 10000
| 90 |  | 10000
| 100 |  | 10000
2) | Table for counts with varying thickness(microns) for 10KeV and
current 10microamperes and gain X20
| Thickness | Counts per
| (microns) | second
| 10 | 1400
| 0.4 | 2300
| 0.15 | 2500
| 0.25 | 2400-2600
| 0.1 | 2500-2650
| 0.3 | 2450-2550
Due to fluctuations in data , the data is treated invalid and
we need to check for a different voltage value or gain
3) | Table for counts with varying thickness(nanometres) for 20KeV
and current 10 microamperes and gain X20
| Thickness | Counts per
| (nm) | second
| 300 | 1380
| 100 | 1350
| 250 | 1325
| 550 | 1380
Here, counts are increasing with thickness, which should not be the case
,so we need to choose a different gain
4) | The counts for thickness 100 and 300 nanometres for 20KeV and
current 10 microamperes and gain X200
| Thickness | Counts per
| (nm) | second
| 100 | 515
| 300 | 510
For a large change is thickness, there is hardly any change in the number of counts.
So we need to use a different gain
5) | Counts for 30KV,20uA,gain x20 and fine gain 1
Thickness | Counts per
(nm) | second
no material | 8000-10000
300nm | 1350
10,000nm | 95 | 10microns
500 | 8000-9000
So the count rate is decaying with a large change in thickness but based on this data
it remained unchanged when we change the thickness by 100 or 200nm
This statement has been verified finding the count rate for different current, voltage and gain
Verification-
1) | SCA TC 346,fine gain-1
Count rates for different thickness, Voltage and Current and Gain
| Thickness/Input Voltage / Current and Gain | 30KV/20uA/x20 | 20KV/20uA/x20 | 10KV/20uA/x20
| 100nm | 8000-9000 | 3000 | 1050
| 300nm | 8000-9000 | 3000-3500 | 1000
| 500nm | 8000-9000 | 3000 | 1000
| Thickness/Input Voltage , Current and Gain | 10KV/50uA/x20 | 10KV/50uA/x50 | 10KV/50uA/x100
| 100nm | 3000 | 30000 | 100000-105000
| 300nm | 3000 | 30000 | 100000
| 500nm | 3000 | 30000 | 100000
| Thickness/Input Voltage , Current and Gain | 10KV/50uA/x200 | 10KV/50uA/x500 | 20KV/10uA/x500
| 100nm | 110000 | 92500-95000 | 2700-3400
| 300nm | 110000 | 100000 | 3200-3600
| 500nm | 110000 | 95000 | 3300-4000
2)SCA AS200,coarse gain X256
| Thickness/Input Voltage , Current | 10KV/10uA | 10KV/20uA
| 100nm | 550 | 1000
| 300nm | 550-600 | 1000
| 500nm | 550 | 1050
6)When we provided the amplified signal from PX2cT(gain 1.5) to SCA instead of directly
connecting the signal from detector to the SCA, a background noise
was observed in the range of 10^4 to 2x10^4. For 500nm,20uA and 10KV,
we get count rate in the range of 4000-5000. The count rate obtained without the amplifier is 1000
This clearly makes it difficult to remove background noise as we cannot subtract it directly

## CAUSES OF ERROR AND CONCLUSION
CAUSES OF ERROR-
1) | Pileup of electron hole pairs causing it to provide a larger signal fluctuating the counts present
2) | Smaller distance between the detector aperture and the gun
3) | Fluctuations in counts per second for larger ranges makes it difficult to select the actual count value
4) | Electronic noise present in the single channel analyzer
5) | Non-Uniformity of the sample
CONCLUSION-
The method is not sufficient for materials having thickness in nanometer
We need to try mitigating the error or employ a better method

## THEORY
THEORY-
If an X-ray photon travels through a material , it either gets scattered , or it’s energy gets absorbed by an electron
, due to which it escapes the atom.
Beer-Lambert’s law states that The intensity of a monochromatic X-ray beam
decreases exponentially with the distance it has traveled inside a material.
I(x) = I0 exp(−µx)
where I0 is the intensity of the initial beam and µ is the attenuation coefficient .
Intensity, which refers to the number of x-rays produced per second,can be measured in terms of
counts per second for sheet of copper with varying thickness.
By measuring the intensities at respective thickness of copper material placed on an aperture we can plot
and verify that Beer Lambert Law is satisfied from which attenuation coefficient can be obtained .